Elemental Analysis by ICP-OES

You are a technician working for a lab which provides analytical chemistry services. Recently a dog hoarding situation was discovered in a rural area of your province.  The dogs were found to be abused and neglected; and poisoning is suspected as dogs displayed symptoms and large amounts of rat poison containing zinc was found in the vicinity of the feeding area.  Samples of the food as well as a brand comparison were given to you to analyze by ICP-OES.  Water samples from the dog’s dishes as well as the home water supply were also given to determine the delivery source.  A sample of a sick dog’s blood was also provided.  You will analyze these samples in an effort to support whether the dogs poor health is indeed a consequence of poisoning.

Objectives:

The objectives of this experiment are to:

- Develop and apply a method for the ICP-OES sample analysis

- Optimize the instrumental conditions for the analysis

- Use a standard sample prep digestion method to prepare dry samples\

- Determine if the poison can be confirmed by elevated levels of zinc provided via food or water

Pre-Lab Questions:

Questions to be answered before starting the experiment:

- In a few sentences, describe the theory behind optical (atomic/ionic) emission spectroscopy (OES).

- How is each element identified in OES? Why it possible to analyse for several elements simultaneously as compared to single analysis as with atomic absorption?

- What considerations must be taken when selecting a line for an element in analysis of a sample mixture? What is the difference between hard and soft lines?

- Describe the ideas behind direct comparison and standard addition methods in the analysis of ICP-OES?  Draw and explain typical calibration curves you would determine during your analysis.

# materials:

Equipment:

- Varian 725ES ICP-OES with a peristaltic pump, autosampler, computer and software; analytical balance, automatic pipettes (10-1000 µL), disposable pipette tips, various glassware.

Reagents and Standards:

- Trace element grade, concentrated Nitric Acid, stock standards 1000 mg/L, Milli-Q water, dilution/wash solution: [2% HNO3 (v/v),  4 L].  The 2% HNO3 is also your standard blank sample.

- Multi-element calibration standard: A multi-element calibration standard stock solution to be made by students.  This is to be made from commercial stock solutions.

- Unknown sample(s) to be determined.

Week 1 : Instrument optimization and Sample preparation:

The steps in this experiment aren’t necessarily performed in order.  If you finish everything in week 1, you can start preparing your dilutions for next week and store them in the fridge.  Also, you may not have time to prepare your sample dilutions after they are finished digesting.

- Sample Digestion and Preparation

You will be working with concentrated acid and Nitrous Oxide fumes.  WEAR GLOVES and perform all work during digestion in the fume hood. MAKE SURE YOU USE THE TRACE GRADE CONCENTRATED NITRIC ACID

- Weigh out 2 seperate 100 mg samples of the brand matched dogfood into Digiprep plastic digestion tubes.  Be sure to label the bottom, side and lid of the tube as heat and acid can sometimes remove even permanent marker.

- Select 2 scene dogfood sample (previously weighed) from box on the bench.

- Place the tubes into the digiprep digestion system located in the fume hood

- Add 3 mLs of concentrated nitric acid. Place the cap loosely on tubes; so that reflux will return to the tube but the fumes can escape.

- Digest the samples for 90 min @ 900C.  Refer to the instruction manual for the digiprep if necessary.

While the samples are digesting, you can move on to the other parts of the experiment.

- Periodically vent the tubes to remove fumes by lifting the lid.  Be careful not to spill any liquid which may have accumulated in the lid.

- After the 90 min is up; remove the lids so the samples will cool quicker.  SLOWLY add 1 mL of 30% H202.  If added to quick, the tubes may foam over.

At this point the samples can be stored in the fridge until next week if time does not permit the completion of the following steps.

- Once cooled; transfer contents of tube into a 15 mL conical centrifuge tube.

- Spin the samples in the centrifuge for 10 min @1500 x g to pellet solid material.  Remove the supernatant and place into labelled 50 mL volumetric flasks bring to volume WITH DDH20. THIS IS THE ONLY TIME WE WILL USE DDH20 AS DILUENT.  THIS IS BECAUSE WE NEED A 2% NITIC ACID CONCENTRATION IN ALL MEASURED SAMPLES.  THE OF THE DILUTIONS WILL BE IN ICP TUBES WITH 2%.

- Multi Element Stock Preparation:

MAKE SURE YOU CHECK ALL LABELS ON THE COMMERCIAL STANDARDS – THEY ARE NOT ALL THE SAME CONCENTRATION

- Prepare a 50 mls of a 10x stock solution in a volumetric flask with the following concentration of minerals from the commercial standards provided. The matrix (diluent) is 2% HNO3 (conc HN03 is 70%).  If you have to make the 2% acid; make sure it is the trace grade Nitric Acid that you use, found in the fume hood.

Copper and Iron – 50 ppm

Magnesium - 250 ppm

Manganese and Zinc – 10 ppm

- Prepare 25 mLs of a 1:5 dilution of the 10x stock solution to be used for optimization using 2% HN03

C: Optimization of Instrument Conditions

It is important to determine optimum operating conditions for each analysis depending upon the element (wavelength/line) of interest.  Emission lines free of spectral interference for the element of interest should be choosen. Other important parameters which affect instrument optimization include RF power, viewing height and nebulizer gas flow.  An explanation of the theory is available from the HELP menu and there is a hard copy next to the computer.  It is advisable to read through this before performing the optimization to fully understand the impact of each parameter.

With this instrument; there are two ways to perform optimization.  The first is manually; in which the parameters are changed by the operator as a solution is pumped through the instrument.  The instrument provides a spectra and the user determines the best settings. Alternatively the software includes a program which performs an automatic optimization based on the range and parameters that you specify. It will increment the settings and then provide optimized values once it has cycled through the range that you specified.

Before performing the optimization a new worksheet must be created and lines selected, which is part of developing your method.

Create New Worksheet

- Open the ICP Expert II software. Under the file menu click New Worksheet in the dialog box.

- Click on the “C” drive, open ICP data/student data/3590.  Name the worksheet and include your group number and date in the title.

- Turn on the argon tank, setting the low pressure gauge to 90 p.s.i.  The system will begin its purge. This takes about 20 to 30 minutes.

Configure tubing and set up Instrument

- Secure the pump tubing:  Select either pump tubing, and place the bottom and top tab between the bottom and top holders stretching the tubing clockwise over the pump rollers.  Repeat for the other tube.  Note that the pump will rotate clockwise, so make sure the tubing is configured in the way so that the tubing will pump the sample into the nebulizer and waste from the spray chamber into to the waste container.

- Lock the pressure bars into place by raising the pressure bar onto the tubing and clamp them to their original position. The middle groove of the 3-channel peristaltic pump is reserved for an internal standard or ionization buffer solution and is not used in this experiment.

- Place the tubing on autosampler pump and clamp pressure bar.

- Once the instrument is finished purging turn on the water bath (under counter behind instrument).

- Light the plasma torch by clicking on the plasma icon at the top of the screen.  (this should not be done untill about 10 min before you are ready to make meaurements.)

Double check with an instructor/TA prior to lighting the torch to make sure the configuration is correct.

Select Lines:

- Click the Methods tab and then the edit method button. In the Method Editor window under elements tab a periodic table is displayed from which lines can be selected.

- Click on an element in the periodic table to display possible line choices.   Click a wavelength and a graph will appear on the top showing the potential interferences (it is also shown in a table on the right side of the screen).  Choose the wavelength for each element based on:

- Intensity (highest signal to noise ratio on the chart to the left side)

- Least interferences (chart on the right side) from elements suspected to be present in your sample

To select the line(s), simply move the cursor to the line and click on OK.  When done click on Close and update your method.

Optimizing the Operating Conditions using Automax:

- Disconnect the sample tubing from the autosampler at the yellow connector and place the instrument end of tubing into a beaker containing your 1 in 5 solution of stock standard.

- In the Method Editor window; click the Conditions tab. At the top centre of the screen it should state that ‘All lines share same conditions’.  If not, go to Options tab in the tool bar menu and check it.

- From the Tools menu in the Method Editor window select automax.

- Choose net signal and check the lines you want included in the optimization.  The default range of values for the parameters is generally acceptable.

- Once finished; the instrument will provide the optimized settings and ask if you want to update the method. Choose yes, but make sure you write down the settings as the update function doesn’t always work.  Before running; verify under the conditions tab that the settings are correct.

ICP-OES Experiment week 2:

- Standard Curve:

- Prepare a standard curve of 8 0.5x serial dilution in the ICP tubes.  You will need a final total volume of 5 mLs left in the ICP tube for analysis.  The diluent Is 2% HN03

  - Label tubes 1-8 (highest to lowest)

  - Aliquot 5 mLs diluent into tubes # 2-8

  - Aliquot 9 mls diluent into tube #1

  - Aliquot 1 mL 10x stock into tube #1

  - Using the pipette; mix tube #1 by aspirating up and down 10x.

  - Remove 5 mLs from tube #1 and add to tube #2.  Mix as in step e.

  - Repeat step f untill tube 8 is reached.

- Finish Sample Prep:

- Prepare a dilutions series of the scene samples (suspected to be high in Zinc) at 1 in 100, 1 in 500, and 1 in 1000.  The easiest and most accurate way is to prepare serial dilutions.   A sample from the 50 mL volumetric should also be analyzed (undiluted).  (for the low abundance elements to compare to the control.)

- Prepare samples of the control dog food from the 50 mL volumetrics – do not dilute!

- Dilute the dog blood sample to 1 in 10, 1 in 50 and 1 in 100 into ICP tubes using 2% HN03

- Dilute the only the dog reservoir  sample 1 in 50 and 1 in 100.  Analyze the house tap water sample straight.

- Place the tubes in the sample rack starting at position 3 (two tubes of only 2% nitric will be the first two samples.) Note that the rack runs from right to left which is opposite the standards rack.

- Complete the Method:

Input concentrations of the External Standards:

- In the Method Editor window click the Standards tab. All lines chosen will display in a form.  Change the number of standards to match what you have.

- Change the units to ppm and enter the assigned standard concentrations values that you have prepared. #1 should be the least concentrated standard.

- The lower part of the screen contains curve analysis options.  Change the % error to 100%.  These options can be manipulated after the data acquisition if necessary.

Defining the Sample Sequence

- Click the Sequence tab to open Worksheet samples table.

- Click the Sequence Editor tab on the right side of the screen which will open a new page. Type in the sample count (remember to add two for the blanks betweem the standards and the samples).

- Check Begin with calibration.  Uncheck the box that says calibrate every 10 samples.

- On the sequence page, click Autosampler Setup tab which will open the Autosampler Setup page. Make sure the Platen Type is SPS 3, Rack Type is Type 60, use is Sample, and Starting tube is 1.  Close the autosampler set up page.

- In the Worksheet Samples table (still under sequence editor) , enter in the sample ID. The first two samples should be blanks. The calibration blank and the standards will already be shown and are not able to be edited here.

Analysis

- Place the calibration blank and standards in the first rack.  (lowest to highest) and the samples in the second rack according to the sequence sample worksheet.

- If necessary; configure tubing and set up instrument as described previously in week 1.  Double check with instructor to verify correct setup.

- Under the Analysis menu click on the green play button at the top of the screen. This will start the analysis, beginning with the calibration blank and standards, and then the samples as defined in the sample sequence.

- Verify the calibration curves for each wavelength. Note that some wavelengths may have failed the calibration for various reasons.  There may also be flags associated with your sample values, for example if a sample is out of the range of the standard; it will not be calculated.  These can be adjusted; consult TA/instructor if help is required.

.

Turning Off the System:

- Rinse the spray chamber by running the wash solution through the system for about 5 minutes.  Run 2% HN03 water for 5 minutes.\

- Click the Plasma off icon. The torch will extinguish and the peristaltic pump will stop.

- Release the pressure bars and lift the tubes out of the grooves on the pump. Do the same with the pump on the autosampler.

- Turn off the water cooler. Wait 5-10 minutes before turning off argon gas.

- Close the valves to the Argon tank and relieve the pressure regulator. Leave the main power switch on to keep the polychromator thermostating system operational.

Data:

The easiest way to collect your data is to copy the values for the standards and samples to “Excel” program and  e-mail/USB transfer the results.  You will need to prepare your own standard curve from the intensity values using a graphing program such as excel.

# Questions:

1. The instrument you used was a radial instrument. The ICP-OES can also have an axial torch.  What are some disadvantages and/or advantages of an axial versus a radial torch.

2. We did not use an internal standard in this analysis.   What would be the advantage of doing so?  What would you have to consider before choosing one?

3. What other detection system could be attached to an ICP and what advantage would there be?